Contents

Preface	2

Ⅰ.	Principle of oxygen measurement with zirconia	2

Ⅱ.	Specification of zirconia oxygen analyzer	4

Ⅲ.	Composition of zirconia oxygen analyzer	4

1	Zirconia detector	4

2	Zirconia converter	6

3	Gas device	7

4	Transformer	8

Ⅳ.	Installation	8

1	Converter	8

2	Installation of stove body flange.	10

3	Installation of detector, diversion tube and dustproof tube	10

4	Selecting position and type for installation	11

Ⅴ.	System adjustment and trial	16

5.1	Connection diagram of CY-2YB disk type instrument system	16

5.2	Connection diagram of CY-2YB hanging type instrument system	17

5.3	System adjustment	17

Ⅵ.	Maintenance and repair	21

Appendix 1. Thermocouple reference table	23

Appendix 2. Contrast table of oxygen electric potential and oxygen content	26

Appendix 3. Contrast table of output signal	27

# Preface

CY series Zirconia Oxygen Analyzer applies to continuous monitor of oxygen content in gas of industrial furnaces and ovens, as basis of adjusting proportion of gas-to-air, or connection with automatic control system to realize hypoxic rational combustion to achieve the goal of cutting down fuel consumption, stabilizing technology, improving quality, reducing environmental pollution and so on. There is clear economic benefit and social benefit.

CY series Zirconia Oxygen Analyzer detector adopts Japanese ion plating film technology so as greatly to prolong service life of zirconia probe, with average service life 18 months, normally used for 2 to 3 years. Sensor adopts the newest sintering technology to rectify big dispersion and poor heat quake property of domestic similar products. The whole reliability and stability of zirconia probe are both in leading position in home.

This instrument converter adopts 16 bit ATMEL series single-chip microprocessor with strong operation capacity, temperature control of zirconia probe within ±2ºC and measure precision of system within ±2 percent. Small signal processing and instrument power supply adopt multi-blocking circuit to isolate various interferences in industrial environment, so instrument runs more reliably. Based on the premise of ensuring measure precision, advanced 3-point calibration reduces maintenance working load of user. Binodal switching output more conveniently meets different demands of users.

# Ⅰ.	Principle of oxygen measurement with zirconia

Zirconia material is a kind of zirconia solid electrolyte, which is stable zirconia ceram sinter formed by high temperature sintering of pure zirconia with certain calcia or yttrium oxide. There are oxygen ion holes in its cubic lattice, so in high temperature condition it is a good oxygen ion conductor.

Fig.1 Concentration cell

Fig.2  Block diagram of zirconia probe detection

By this specific property, in certain temperature, when there are different oxygen contents at both sides of sensor, it is a typical oxygen concentration cell. If pure platinum electrode is coated on inside and outside of zirconia tube, the zirconia tube is heated by electric stove, and its inner and outer wall contact gas with different oxygen contents, the zirconia tube is an oxygen concentration cell. The following reaction will occur on both platinum electrodes.

For electrode at air side (reference gas side): O2+4e→2O2-

For electrode at low oxygen side (measured side): 2O2 - →O2+4e

After the two kinds of migration reach balance, an electric potential signal E related to oxygen concentration occurs between electrodes.

Oxygen electric potential E complies with Nernst equation:

Where: R- gas constant,

T- absolute temperature of zirconia tube,

F- Faraday constant,

PX- oxygen content percent of measured gas,

PA- oxygen content percent of reference gas, normally 20.6%.

# Ⅱ.	Specification of zirconia oxygen analyzer

* 	Measured object: concentration of gas or mixed gas of various industrial furnaces and ovens

* 	Measure element: zirconia tube

1. 	Oxygen measure range: 0~20.6% O2  or 0~10% O2

2. 	Precision of the instrument: intrinsic error of system measure within ±2% full range

3. 	Precision of transmitter: 1.0 class (within 1.0% full range)

4. 	Precision of temperature control: constant temperature point 700±2ºC

5. 	Response time: ≤ 3 sec, (reaching 90% response)

6. 	Alarming output: upper and lower limit outputs, “normally open” or “normally close” point optional

7. 	Analogue output signal: 4~20mA ADC (load 0Ω~750Ω) corresponding to oxygen content 0~10% O2 or 0~20.6% O2

8. 	Background correction range: -20 mV ~ + 20 mV

9. 	Digit display mode: LED four nixie light

10. 	Power supply: AC220V±15%

11. 	Power consumption: < 6W (non-including heating power)

12. 	Heating power: about 50W (average), able to supply 150W power output

13. 	Environmental condition: temperature 0~50ºC

Relative humidity < 90%

14. 	Weight: detector about 10 kg, converter about 5 kg

# Ⅲ.	Composition of zirconia oxygen analyzer

This instrument consists of four portions of zirconia detector (probe), zirconia converter, transformer and gas device.

## 1	Zirconia detector

For structure of zirconia detector, see Fig.3. Oxygen sensor is installed at the head of detector and fixed by four stainless steel bolts. There are two gas ways, standard gas inlet and reference gas inlet, in detector. Different standard gas samples enter the standard gas inlet, and the system is on-line calibrated by converter. The “standard gas inlet” in the figure is used for calibrating operation. If no calibration is made, please plug up this “standard gas inlet” to avoid air from entering to bring observation error. The “reference gas inlet” should be open to air at all time.

Fig. 3 Structure of zirconia detector

(Taking detector with standard length 400 mm as an example)

There are 8 leading terminals in junction box of detector, see Fig.4. There is temperature sensor (AD590) in detector, so when thermocouple signal is transmitted from detector to converter, compensation conducting wire need not be used.

There are six kinds of standard length for detector, 180mm, 400mm, 600mm, 800mm, 1000mm and 1200mm.

Fig.4  Detector terminals

## 2	Zirconia converter

2.1 	According to installing mode, CY series zirconia converter is divided into two models, CY-2YB (disk type) and CY-2YB (hanging type), where for disk type there are two kinds of specification optional.

The converter adopts 16 bit ATMEL single-chip micro computer system to process data with strong function and high precision. Advanced menu operation management is adopted so as to be convenient and fast to look up data and calibrate on-line. Because 3-point calibration mode is adopted, and condition points can be corrected and calibrated at any time, so as to reduce maintenance work intensity of operators and maintainers. Besides accurate 4~20 mA output, converter also provides optional upper/lower limit output or 485 communication output for user to operate simply.

2.2 	Disk type converter (CY-2YB)

Rear terminal of CY-2YB zirconia converter

Fig.5  CY-2YB converter size

2.3 	Hanging type converter (CY-2YB)

Fig.6  Terminal of CY-2YB (hanging type) zirconia converter

Fig.7  CY-2YB converter size

## 3	Gas device

Gas device may be divided into dustproof tube, diversion tube, dust cover, metal filter and high temperature gas device.

## 4	Transformer

In order to prolong the service life of ISG-C zirconia probe, the instrument is equipped with a separate transformer T100, see Fig. 8.

4.1 	T100 transformer

Fig.8  Overall size of T100 transformer

# Ⅳ.	Installation

## 1	Converter

1.1	Installed portions of converter should meet the following requirements

* 	To avoid mechanical vibration as far as possible

* 	To avoid high humidity condition

* 	Environmental temperature within standard range

* 	To avoid strong magnetic field interference

* 	To avoid interference of power line, motor, exciting relay, pump and so on.

* 	To provide enough maintenance room

* 	Converter installed as close as possible to detector with the distance no more than 200m

1.2 	Installation of disk type converter (CY-2YB)

According to the arrangement of instrument cabinet, select horizontal or vertical mounting holes on panel of the cabinet. Size of the open holes in the panel of the cabinet to mount the converter is 75+1×150+1

1.3	Installation of hanging type converter (CY-2YB)

Fig.9 Installation diagram of CY-2YB converter

* 	Watch case installation: take out fixing boards from both sides of the converter, insert it into the panel of the cabinet, the size of open holes per the figure, and fasten it with the fixing boards.

* 	Wall installation: turn the fixing boards upside down and fix the converter with the bottom of converter shell against wall.

## 2	Installation of stove body flange: According to Fig. 10, make an opening with diameter no less than 110mm in stove wall, first weld built-in stove wall buried tube. When installing oxygen analyzer, stove body flange (supplied by manufacturer) is welded on the stove wall buried tube. There L>20, D>110 and K=90 in the figure.

Fig.10 Stove wall buried tube and stove body flange

## 3	Installation of detector, diversion tube and dustproof tube

Before installation, prepare M12×50 bolts, corresponding nuts, filler sheets, gaskets and seal ring in the flange and connect tightly with stove body flange (seal ring, bolts and so on supplied by manufacturer).

If dust cover or metal filter is used, dust cover or metal filter has been installed on the top of detector before leaving factory.

Notice: During installation, do avoid knocking, bumping or mechanical vibrating. Bolts must be fastened to ensure leakproofness.

## 4	Selecting position and type for installation

4.1	Selecting position for installation

Select such position to make gas flow easy and typical gas sample taken to check. Do not install the instrument in the inner corner, whirling corner or neck portion of stove to avoid oxygen content fluctuation, gathering dirt much or responding slowly. Taking boiler of power plant as an example, installing position should be at inlet/outlet of economizer or air pre-heater.

4.2	For selection of detector length following methods may be adopted.

* 	Length of direct inserting type: over 300 mm + thickness of stove wall, i.e. the front end of probe extends 300mm from stove wall.

* 	Diversion tube type: length of diversion tube is more than 300 mm + thickness of stove wall, and length of probe should be no less than thickness of stove wall

The installing method of diversion tube type is recommended. This method can improve use effect of zirconia probe and prolong its service life.

4.3	Introducing several sampling methods as follows

* 	Direct inserting type, for high dust occasion, dustproof tube should be put over detector, or use dust cover directly, see Fig.15.

4(a) Horizontal installation                  4(b) Vertical installation

Fig.11  Dustproof installation

* 	Diversion tube type: standard length of detector is 180 mm, 400 mm, 600 mm, 800 mm, 1000 mm and 1200 mm; because of thicker stove wall and shorter detector, sampling distance should be prolonged, equipped with ZDL-L diversion tube. For installation, see Fig. 12.

Fig.12 Diversion tube type installation

* 	High-temperature type: Gas device is divided into draught type (see Fig.13) and extraction type (see Fig.14). The latter is used for the occasion with negative pressure more than -50Pa, equipped with jet pump. Gauge range of air source used for jet pump is from 0.05MPa to 0.15MPa. Jet pump is adjusted with air filter regulating valve.

Fig.13  High-temperature draught type installation

Fig.14  High-temperature extraction type installation

4.4 	Wiring and pipe fitting

4.4.1 	When selecting power cable and signal cable for the instrument, consider environmental temperature of the installed place, resistance of conducting wire, other signal source, and wiring path between detector and converter. For cable to be used, refer to the following table.

|  | Below 60ºC | 60~80º C | 80~150ºC |
| --- | --- | --- | --- |
| Signal cable (1)oxygen cell signal (2)thermocouple signal (3)compensation signal | Six-lead cable, with 1 square mm cross section of each lead, with polyvinyl chloride insulation protective layer | Six-lead cable, with 1 square mm cross section of each lead, with heat resistant polyvinyl chloride insulation protective layer | Six-lead cable, with 1 square mm cross section of each lead, with 600V class silicone rubber and glass fiber protective layer |
| Heating furnace cable | Two cables each with 1 square mm cross section | Two cables each with 1 square mm cross section | Two cables each with 1 square mm cross section |

For signal cable, it would be best to use shield cable with shield lay grounded. If signal cable is not shield cable, it should be installed as far as possible away from the wires of process heater or the power cable of large motor.

4.4.2	Cable channel

It is recommended to use tube with 22 mm outer diameter or flexible conduit with 19 mm outer diameter as cable channel. Power cable channel or ground channel should be away from signal cable channel to avoid signal interference.

4.5 	Complete set and spare parts table

| No | Description | Model and type | Unit | Specification |
| --- | --- | --- | --- | --- |
| A01 | Zirconia analyzer | Low- temperature A model | 1 set | Detector: 180mm, 400mm, 600mm, gas resistant temperature < 600ºC Converter: A: CE-2D, output 4-20mA. B: CE-2C, output 4-20mA. Showing range:0-20.6%. Heating power 80W (max 150W). Dustproof tube: insert length 250mm, 500mm, 700mm |
|  |  |  |  | See installation diagram 11 |
| A02 | Zirconia analyzer | Low- temperature B model | 1 set | Detector, converter, transformer are the same as the above (adding dust cover or filter): 800mm, 1m, 1.2m |
| A03 | Zirconia analyzer | Middle- temperature M model | 1 set | Detector, converter, transformer are the same as the above, diversion tube: 800mm, 1m, 1.2m, gas resistant temperature < 900ºC |
|  |  |  |  | See installation diagram 12 |
| A04 | Zirconia analyzer | High- temperature H model | 1 set | Detector, converter, transformer are the same as the above, gas device: three-way + carbon silicide tube, gas resistant temperature 900-1400ºC, insert length: 800mm, 1m, 1.2m Jet pump: input pressure 0.05 - 0.15 MPa (obtained through reduce valve), output pressure 20 – 200 mm water column |
|  |  |  |  | See installation diagram 13 and 14 |
| A05 | Detector | ISG-C | 1 piece | 180mm, 400mm, 600mm, 1000mm, 1200mm |
| A06 | Zirconia tube assembly | CSG-2 | 1 set | Zirconia tube, seal ring, filter net, bolts, diversion tube |
| A07 | Converter | CE-2D CE-2C | 1 set 1set | Oxygen content: 0-10%, 0-20.6% Output 4-20mA, upper/lower limit output (optional) 485 output (optional) |
| A08 | Transformer | T100 | 1 set | Input 220VAC, output 90-110VAC |
| A09 | Low-temperature dustproof tube | DFC-L | 1 piece | Gas resistant temperature ≤ 600ºC, length 250 mm, 500 mm, 700 mm |
| A10 | Low-temperature diversion tube | DDL-L | 1 piece | Gas resistant temperature ≤ 600ºC, length 800mm, 1m, 1.2m |
| A11 | Middle-temperature diversion tube | ZDL-L | 1 piece | Gas resistant temperature ≤ 900ºC, length 800mm, 1m, 1.2m |
| A12 | High-temperature diversion tube | GDL-L | 1 piece | Gas resistant temperature ≤ 1400ºC, length 800mm, 1m, 1.2m |
| A13 | Correcting air source | XV-1 | 1 set | User must be fitted with correcting air source, multiple instruments may be equipped with one source |
| A14 | Stove body flange tube |  | 1 set |  |
| A15 | Jet pump | PSB-1 | 1 set |  |
| A16 | Technical service | Manufacturer supplies site service of adjusting, trial, setting and running for instrument, with service charge |

# Ⅴ.	System adjustment and trial

## 5.1	Connection diagram of CY-2YB disk type instrument system

Disk type instrument system consists of ISG-C detector (probe), CY-2D converter, T100 transformer and connection wiring, see Fig. 15.

ISG-C zirconia probe terminals

Fig.15   Connection diagram of CY-2YB (disk type) zirconia analyzer system

## 5.2	Connection diagram of CY-2YB hanging type instrument system

Hanging type instrument system consists of ISG-C detector (probe), CY-2YB converter, T100 transformer and connection wiring, see Fig. 16.

CY-2YB converter terminals

ISG-C zirconia probe terminals

Fig.16  Connection diagram of CY-2YB (hanging type) zirconia analyzer system

## 5.3	System adjustment

5.3.1	Function of panel key

For panel of CY series zirconia converter, see Fig.17.

Panel of CY-2YB (disk type) zirconia converter

Fig.17  Panel of CY-2YB (hanging type) zirconia converter

* 	Display: consists of 4 nixie lights and 8 LED indicator lights.

Nixie lights are used to show oxygen content, temperature, oxygen electric potential and other parameters.

LED indicator lights show current work condition including “oxygen content”, “oxygen electric potential”, ”temperature”, ”correct”, ”run”, “communication”, “upper limit”, and “lower limit”.

* 	Panel keys: including “enter” key, “return” key, and “▲” key, “▼” key,  “   ” key.

* 	Enter key: press this key to enter adjusting and correcting menu or confirm calibration value.

* 	Return key: press this key, operation returns to upper menu.

*	▲key: press this key, perform “plus” operation on current submenu or value.

*	▼ key: press this key, perform “minus” operation on current submenu or value.

*	   key: press this key, perform “plus” operation on current main menu. This key also has “shift” function. Through this key, select figure place to be adjusted.

5.3.2 	Operation instruction and adjusting method

If press any key for more than 10 seconds, the system will restart.

Run

1. 	When zirconia temperature is lower than 600 degrees, zirconia temperature is shown.

2. 	When zirconia temperature is higher than 600 degrees, oxygen content is shown.

3. 	In normal working condition, if press “▲” and “▼” keys, “oxygen content” and “zirconia temperature” displays are switched. Zirconia temperature shows for over 2 seconds, the instrument will automatically switch to “oxygen content” display.

Calibration

1. 	Press “enter” key, input code “0001” with  “▲”, “▼”, “   ” keys, press  “enter” key again to enter calibration interface. The instrument displays as follows.

| 1 | - | - | 1 |
| --- | --- | --- | --- |

Main menu       submenu

2. 	Press “  “ key to make the instrument show the following interface:

| 1 | - | - | 1 |
| --- | --- | --- | --- |

Main menu       submenu

3. 	This interface is on-line interface of oxygen content. Perform oxygen content air calibration (the instrument shows     ), oxygen content range calibration (the instrument shows       ), and oxygen content user condition point calibration (the instrument shows       ).

In its flashing condition, press “▲” key to select submenu. “      “, “      “, “      “ separately are shown in menu. Press “enter” to enter the condition confirmed.

Oxygen content calibration (note: ensure the temperature of zirconia probe within 700±2ºC.)

oxygen content air calibration [i.e. correcting zero electric potential]: Open “standard gas inlet” of detector to make air enter. Wait for stability. Then press “enter” to confirm. 20.6 is shown.

oxygen content range calibration: Open “standard gas inlet” of detector to make 0.5 – 2.5% standard gas enter. After stability, input the value of standard gas. Press “enter” to confirm. See Fig.18.

oxygen content user condition point calibration (3% - 18%): Confirm the instrument  runs normally. After stability, input the value of oxygen content in correct condition. Press “enter” to confirm. See Fig.18.

Fig.18  Connection diagram of standard gas and detector

Note: Output flow of standard gas should be no more than 1 liter per minute.

Note: If calibrate only one point, the other two points still retain the last calibrated values. If calibrate two points, the converter will automatically produce operational slope on the two points, and the last calibrated value of the other point will produce small deviation. It is recommended that when calibrate two points, use as far as possible standard gas similar to your condition for the calibration of oxygen content user condition point calibration (      ). After calibration, do press “return” key to return to running condition.

4. 	When there is error in system connection or signal input, CY-2YB or CY-2YB converter automatically shows a series of error codes to indicate current error condition.

Error code set

EE01  electric stove wire short in sensor

EE02  electric stove wire open in sensor

EE02  non-matched sensor (made by non-original factory)

E000  calibration menu code input error

E101  temperature zero calibration input signal error or signal not connected yet

E102  temperature range calibration input signal error or signal not connected yet

E201  oxygen electric potential zero calibration input signal error or signal not connected yet

E202  oxygen electric potential range calibration input signal error or signal not connected yet

E301  oxygen content zero calibration input signal error or signal not connected yet

E302  oxygen content range calibration input signal error or signal not connected yet

E303  oxygen content user any point calibration input signal error or signal not connected yet

# Ⅵ.	Maintenance and repair

In order to make the instrument in good running condition, the instrument must be maintained and repaired, according to the special installed environment, preventive maintenance should be made monthly or at least once a year.

6.1 	Maintenance item

6.1.1	According to operating method on the panel of converter, check every function.

6.1.2	Check the filter at the top of detector or the inside of gas device tube for dirt. If too dirt, clean it away, otherwise measure precision and response time will be affected.

6.2 	Common problems

6.2.1	Working temperature of zirconia tube lower than 700ºC

| Cause | Solution |
| --- | --- |
| Broken fuse | Replace fuse |
| Broken leading wire of electric stove | Repair or replace |
| Broken electric stove wire | Resistance of electric stove should be about 80Ω. If electric stove wire is broken, it should go back to the factory to repair. |
| Faulty thermocouple | Normally, resistance of thermocouple wire is no more than 5Ω. If abnormal, it should go back to the factory to repair. |
| Broken thermocouple wire | Confirm thermocouple wire is broken (not faulty contact). It should go back to the factory to repair. |
| Temperature rises too slowly, reaching 700ºC within 1 hour | When temperature is tens of degrees lower than 700ºC, maybe electric power of the electric stove is not enough. Adjust output connection of transformer to improve output voltage. |
| Thermocouple wire connection in an opposite direction | Check thermocouple connection. Exchange “+” and “-“ connection. |

6.2.2	Working temperature of zirconia tube higher than 780ºC

Check working voltage of electric stove. If the voltage is always higher than 90 VAC (the output voltage of transformer), that means there is fault in the temperature control circuit of the instrument. First replace double-direction silicon-controlled rectifier.

6.2.3	Faulty temperature compensation circuit

If “E - - -“ is shown, that means faulty temperature compensation circuit. Now first check temperature sensor (AD590) and its leading wire for good contact and correct polarity. Next probably replace temperature sensor (AD590).

6.2.4 	Abnormal oxygen content display

* 	If the following cases occur, i.e. loose screw of standard gas hole, no asbestos pad on the detector, loose mounting screw of detector, disturbance produced by air inleakage will make display of oxygen content abnormal.

* 	When zirconia tube used reaches its service life and the inner resistance of the zirconia tube is more than 800Ω, display of oxygen content will be abnormal.

## Appendix 1. Reference table of nickel-chromium/ nickel-silicon (nickel-chromium/nickel-aluminum) thermocouple

K thermocouple index R

(Temperature of free junction ºC)

| Temperature of working junction ºC | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Millivolt (abvolt) |

| Temperature of working junction ºC | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Millivolt (abvolt) |

| Temperature of working junction ºC | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Millivolt (abvolt) |

## Appendix 2. Contrast table of oxygen electric potential and oxygen content

| E (mV) | Oxygen content, % | E (mV) | Oxygen content, % | E (mV) | Oxygen content, % |
| --- | --- | --- | --- | --- | --- |

## Appendix 3. Contrast table of output signal

| Oxygen content (percentage) | Output signal |
| --- | --- |
|  | 4 – 20 mA |
| 0 3.25 5.15 8.40 10.3 13.55 15.45 18.70 20.60 | 4 6.25 8 10.52 12 14.52 16 18.52 20 |

# Text boxes

Voltmeter

electrode

Gas

Standard gas inlet

Reference gas

Cable hole

Oxygen cell output

Electric stove power supply HTR

Compensation output CJ

Thermocouple output TC

Standard gas input

Reference gas input

Cover silicon

bolt for cover

Reference gas inlet

A-A cross-section

Upper limit

1 Cell 2      3 Couple 4     5 Temperature  6

+    -      +       -     + compensation  -

Electric stove

Zirconia element

End cover fixing hole

Lower limit

Transformer input

HTR Heating output

output

220VAC      Transformer    heating output         Upper limit    Lower limit      4~20mA

input           input         HTR                                            output

Cable hole

Mounting hole

Hole spacing

Hole spacing

L: according to thickness of stove wall

Stove body flange

Stove wall buried tube

Stove wall

Flue gas flow direction

Detector

Dustproof tube

Gas flow direction

Junction box

Detector

Stove wall

Exhaust port

Dust leakage

Diversion tube

Stove wall

Gasket

Stove wall buried tube

Stove body flange

Stove wall

Stove wall buried tube

Stove body flange

High-temperature diversion tube

High-temperature sampler

Gasket

Fuel gas direction

Stove wall

High-temperature sampler

High-temperature diversion tube

Stove wall flange

Compressed air

Exhaust direction

Jet pump

Gasket

Fuel gas direction

Stove wall

Detector

Stove body flange

Stove wall buried tube

Electric stove

Transformer terminals

+Cell- +Thermocouple- +Temperature compensation-

output

HTR heating output

Transformer input

Lower limit

Upper limit

Transformer terminals

Electric stove

+Cell- +Thermocouple- +Temperature compensation-

output

Upper limit      Lower limit

input

Transformer      HTR

input       heating output

Standard gas

Flow adjusting knob

Rubber tube

Standard gas cock

Flow meter

Detector